Note:

- Please read the operation manual carefully before using the product.

Warning:

- Connect the power cord to the wall socket directly, and avoid using the extended electric wire.

- If the power cord or plug had wear and other damage, please disconnect the plug. (Hold the plug instead of the wire)

- If following situations happened, please turn off the power supply and disconnect the plug. (Hold the plug instead of the wire)

1. Fluid splash on the pump.

2. You think the pump need to maintain or repair.

- The user’s power socket must have ground wire, and have reliable grounding.

Note: The foot pedal switch and other external control plugs must be connected or discounted in the power-off status to prevent the external control interface from being burned

Catalogue

1. Product Introduction	- 1 -

2. Product Appearance	- 1 -

3. Keypad Instruction	- 2 -

4. Operation Interface	- 3 -

5. External Control Interface Instruction	- 11 -

6. Technical Specification	- 15 -

7. Main Functions & Features	- 15 -

8. Dimension Drawing	- 16 -

9. Maintenance	- 18 -

10. Warranty and After-sales Service	- 19 -

# Product Introduction

N6-30L is a type of standard industrial peristaltic pump driven by servo motor, with 3.2 inch true color LED screen and rotary knob speed control & pure imported keypad. The speed range is 0.1~300 rpm. It has variety of external control modes for option and supports RS232/RS485 and standard MODBUS communication protocol (RTU mode) to meet industrial sites.

Suitable pump head: DZ45.

# Product Appearance

A——Drive

B——Pump head

C——Power switch

D——Power socket

E——External control interface

F——Internal fuse

# Keypad Instruction

- Start/stop button: Press this button, the pump starts running with current speed. Press this button again, pump stops running.

- Change direction button: Change the running direction of the pump every time you press the button.

- Full speed button: Press this button at stop status or transferring mode, the pump will run with highest speed. This function is used for fast fill liquid into tube or wash the tube. At timing tests, this button is invalid.

- Digital knob: Rotate this button at stop status or transferring mode to adjust speed.

In stop status, press [digital knob] to enter [selecting menu interface], then rotate [digital knob] to choose relative items. Press [digital knob] to change the parameters when setting the parameters.

- ESC button: In menu interface, press this button back to previous menu. In main interface, keep pressing this button for more than 5 seconds to check pump SN code and programme version when the pump was in stops status. And keep pressing this button and turn on the power supply at the same time, that will initialize the pump and all the parameters will be lost.

# Operation Interface

  - Main Interface

  - Menu Setting Interface

Steps:

- Press [digital knob] in the main interface to enter the menu setting interface,       press again to select the setting parameters.

- Rotary the [digital knob] to move up and down to select the parameter to be set.

- Press [digital knob] to enter the next setting parameters interface.

- Press [ESC] to back to main interface.

Note: Other buttons are invalid in this interface.

  - Calibration Interface

Steps:

- After entering this interface, the test time defaults to 60 seconds, and the actual liquid volume defaults to the liquid volume displayed according to the current set flow rate.

- Press the [digital knob] to select the test time. In the (1) state, turn the [knob], cursor moves to the [volume] line, and then press the [Start] button to start the pump for calibration. A countdown is displayed at the test time.

- After the countdown is over, the cursor moves directly to the actual liquid volume, then turn the knob enters the actual liquid volume.

- You can select actual volume, pressing [digital knob] to set actual volume.

- You can also switch to the (2) state by pressing the [digital knob] when in the (1) state, then turn the [digital knob] in the (2) state to set the test time.

- After enter the actual volume, press the [digital knob] to confirm the entry. The ‘Calibration finished!’ dialog box will pop up. Press the knob again back to calibration interface.

- Reset function: Select ‘Yes’ and press[digital knob], the ‘Reset calibration finished!’ dialog box will pop up, the calibrate factor been restored.

  - Select Pump Head & Tube Interface

Steps:

- Select the pump head model by pressing the [digital knob] in Head & Tube interface. In the (1) state, turn the [digital knob] to move the cursor up and down to select set pump head or tube size.

- In the (1) state, press the [digital knob] to switch to the (2) state.

- In the (2) state, turn the [digital knob] to select the pump head model.

- In the (2) state, press the [digital knob] or [ESC] to return to the (1) state and confirm the selected pump head or tube size.

  - Dispensing Interface

Steps:

- Press [digital knob] to select dispensing in the dispensing interface. In (1) state, turn the knob to select setting On/Off, Time or Unit.

- In (1) state, press the knob change the interface to (2) state.

- In (1) state, press [ESC] back to menu interface.

- In (2) state, turn knob to turn On/off dispensing function, or set the dispensing time (0.1-9999s), or set the time unit (second, minute, hour).

- In (2) state, press knob or [ESC] back to (1) state and confirm the setting parameters.

- After turning on the dispensing, the pump will automatically stop when reach the setting time.

  - External Control Speed Interface

Steps:

- Press [digital knob] on the external control speed interface to select the external control speed. In the (1) state, turn the knob to move up and down to select the external control speed, signal type or maximum speed.

- In the (1) state, press knob to the (2) state.

- In the (1) state, press the [ESC] back to menu interface.

- In the (2) state, turn the [digital knob] to select the external control speed control on/off or select the signal type: 0-5V/0-10V/4-20mA or select the maximum speed.

- In the (2) state, press the [digital knob] or [ESC] to return to the (1) state and confirm the selected item.

- In the (1) state, keep rotating the [digital knob] in clockwise to enter the external control speed customization interface, as shown below:

- In the (1) state, turn the [digital knob] to select (0V corresponding speed: 0) or (5V corresponding speed: 300). The relative speed can be set freely within the specified speed range. The pump will automatically create a liner speed relationship between the analog values.

- In the (1) state, press the [digital knob] to the (2) state.

- In the (1) state, press the [ESC] to return to the [Setup Menu Selection Screen].

- In the (2) state, turn the [digital knob] to set the corresponding speed for each part.

- In the (2) state, press the [digital knob] or the [ESC] to return to the (1) state and confirm the selected item.

Select external control signal according to external input signal, 0-5V, 0-10V, 4-20mA for option. External control speed can set the maximum speed, when the maximum speed is 300rpm, the analog signal voltage range has a linear relationship with the motor speed.

If the maximum speed is not 300 rpm, the motor speed will be limited by the analog signal. When the motor speed and the analog signal amount reach the set maximum speed according to the corresponding proportional relationship, then when the analog signal amount is increased, the motor speed is set by The operation at a fixed maximum speed does not follow the increase of analog signal quantity. For example: set 0V to 0 rpm, 5V to 300 rpm (2.5V should be 150 rpm), the maximum speed is set to 150 rpm, if the external input analog signal is 2.5V, then the motor speed is 150 rpm, input signal The motor speed remains unchanged at 150 rpm after exceeding 2.5V.

  - External Control Interface

Operation steps:

- In the (1) state, turn the [digital knob] to select the external control start/stop, change direction or signal type.

- In the (1) state, press the [digital knob] to the (2) state.

- In the (1) state, press the [ESC] to return to the menu interface.

- In the (2) state, turn the [digital knob] to select the external control start/stop on/off or external control direction on/off or select the signal type: level/pulse ( in level mode, the corresponding internal control buttons are invalid.)

- In the (2) state, press the [ESC] or the [digital knob] to back to the (1) state and confirm the selected item.

- External control full-speed function is level mode. Connection is full speed, disconnection means stop full speed;

- External control start/stop, change direction signal is active signal and passive signal.

- External control start/stop, change direction signal: Level mode and pulse mode, detailed interface refer to external control interface instruction.

- After turn on the dispensing function, external control start/stop only valid in pulse mode, invalid in level mode.

- External control have independent switch, only works when the corresponding external control function is turned on.

  - Communication Setting Interface

Operation steps:

Serial communication is always on and the On/Off option is not set.

- In the (1) state, turn the [digital knob] to select setting local address, communication port or baud rate.

- In the (1) state, press the [digital knob] change to the (2) state.

- In the (1) state, press the [ESC] to return to menu interface.

- In the (2) state, turn the [digital knob] to select the local address (01-32) or the interface to select RS232/RS485 or baud rate selection: 2400/ 4800/ 9600/ 19200.

- In the (2) state, press the [ESC] or the [digital knob] to return to the (1) state and confirm the setting parameters.

The pump support Modbus communication protocol----RTU mode, communication port is RS485 or RS232.

Note:

After the setting is completed, the pump only receive communication signal control in main interface; in other setting interface, the communication is invalid.

  - Back Suction Interface

Operation steps:

- In the (1) state, turn the [ digital knob] to select setting on/off or suction angle.

- In the (1) state, press the [ digital knob] change to the (2) state.

- In the (1) state, press the [ESC] to return to menu interface.

- In the (2) state, turn the [digital knob] to set the back suction function on/off or to set the back suction angle (0-360°).

- In the (2) state, press the [ESC] or the [ digital knob] to return to the (1) state and confirm the setting parameters.

Note: When the back suction is off, the main interface does not display the back suction sign. Only when the back suction is turned on, the main interface will display the back suction sign.

# External Control Interface Instruction

External control interface like below picture:

- Analog signal input port: In the external control setting interface to select external control speed signal then turn on the function of external control speed. Control the motor speed from 0rpm to the highest speed through analog signal.

0-10V: 0V to 10V voltage signal input terminal.

0-5V: 0V to 5V voltage signal input terminal.

4-20mA: 4mA to 20mA current signal input terminal.

I_/V_: Analog signal negative terminal.

Note: Please do not connect 0-10V signal with 0-5V terminal or 4-20mA input terminal. This is forbidden. Wrong connection will damage the pump.

- External control signal input terminal

- Internal isolation 5VDC output.

+5V: Internal 5V output positive pole.

GD1: Internal 5V output negative pole.

- External control start/stop, direction, full speed signal input terminal: Active signal input, 5-24VDC.

GD2: External control signal common input terminal.

NC: External control full speed input terminal.

CW/CCW: External control direction signal input.

R/S 2: External control start/stop signal input.

Note: Set up the external control mode in the setting interface, turn on the correspond external control function, external control signal is active.(The external control full speed is default level mode, short-circuit is on, disconnecting is off, it does not need to be opened in the settings interface)

- External control start/stop signal input terminal: Passive signal input

R/S 1: External control start/stop signal.

This terminal can connect with passive switch and foot pedal. In the external control setting interface to set the active of this terminal, external control start/stop (passive) the switch option.

External control wiring:

- In pulse mode, disconnect K2 after short circuited, the motor starts working. Disconnect again after short circuited, the motor stops working.

In level mode, short circuited K2, the motor starts working. Disconnect K2, the motor stops working.

- In pulse mode, every time disconnect the K1 after short circuited, the motor changes working direction.

In level mode, short circuited K1, the motor runs in clockwise. Disconnect K1, the motor runs in counterclockwise.

- In pulse mode, disconnect K3 after short circuited, the motor starts working. Disconnect again after short circuited, the motor stops working.

In level mode, short connect K3, the motor starts working. Disconnect K3, the motor stops working.

- Short circuited K4, the motor will run with full speed, disconnect K4, the motor is restored to the state before full speed operation.(External control full speed defaults to level mode)

The motor working status output terminal: Output motor working status as below:

- Communication port

RS232 Communication: Choose RS232 in the communication setting interface, this terminal is active.

GND: Communication ground terminal.

TXD: Master sends, peristaltic pump receives signal terminal.

RXD: Peristaltic pump sends, master receives signal terminal.

RS232 Communication Interface Connection Diagram as below:

RS485 Communication Interface: Choose RS485 in the communication setting interface, this terminal is active.

GD1: RS485 signal ground

A+: Connect RS485 A+ terminal

- : Connect RS485 B- terminal

Note that: when leaving the factory, the external control interface will be plugged with terminals. If you need to use other external control equipment of our company, such as foot pedal, hand-held dispensers, etc., please unplug the terminals first, and then plug the external control equipment.

# Technical Specification

| Start/stop,direction signal | Passive switch signal, like foot pedal; Active switch signal: 5V-24V | Power supply | AC220V±10%, 50Hz/60Hz(standard) |
| --- | --- | --- | --- |
|  |  |  | AC110V±10%, 50Hz/60Hz(option) |
| Weight | 25.7kg | Power consumption | N6-30L: <300W |
| Speed resolution | 0. 1rpm | Communication interface | RS232, RS485 supports Modbus communication (RTU mode) |
| Operation method | Mechanical keypad and digital rotary knob speed control | Dimension (L*W*H) | N6-30L:299.4*303*278mm |
| Display method | 3.2 inch LCD | External control speed signal | 0-5V, 0-10V, 4-20mA for option |
| Relative humidity | <80% | Speed range | 0.1-300rpm |
| IP rate | IP31 | Temperature | 0-40℃ |

# Main Functions & Features

- 3.2 inch screen LCD display;

- Super silent drive setting, precise control, low vibration and low low noise;

- Mechanical keypad and digital knob control, convenient for users setting the parameters;

- Timing function, the time range is 0.1s-9999 hours; Repeat number setting range is 1-9999 times, setting “0” means unlimited.

- Various external control functions, support 0-5V, 0-10V, 4-20mA analog signal control speed.

- Support RS232/ RS485 communication, standard MODBUS communication protocol (RTU mode)

- Power down memory function, store parameters in time, safe and reliable.

- Strong anti-jamming feature, wide voltage design, suitable for complex power supply environment.

- Fast fluid-filled function, not only can clean the tubing, but also fill liquid into the tubing.

- High torque and low power loss.

# Dimension Drawing

(Unit: mm)

  - Single pump head

DZ45-I Pump Head

DZ45-II Pump Head

  - N6-30L dimension drawing

N6-30L series + DZ45-I pump head

N6-30L series + DZ45-II Pump Head

# Maintenance

- Check the running status of machine before starting it, normal operation can be put into use.

- Check for leakage, and correct fault which can be appeared.

- Clean liquid overflowed from the pump in time.

- Please turn off the power supply and unplug the power socket (Hold the socket instead of power cord) when liquid splashed on pump. Check whether liquid flows into the machine, if it does, please contact the manufacture.

- The foot pedal switch and other external control plugs must be connected or disconnected in the power-off status to prevent the external control interface from being burned.

- The user’s power socket must have ground wire, and have reliable grounding.

- This product has no waterproof measures. Please take protective measures when using in water environment.

- This product does not have special certification such as medical certification. When it needs to be used in special fields such as medical and military, please self-certify.

- If the pump does not use for a long time, please clean it and keep it in dry and ventilated environment.

- The company shall not bear the direct and indirect losses caused by the malfunction or improper operation of this product.

# Warranty and After-sales Service

We support 3 year warranty for the pumps, subject to the exceptions below. Our company shall not be liable for any loss, damage, or expense directly or indirectly related to or arising out of the use of its products.

This warranty does not obligate our company to bear any costs of removal, installation, transportation, or other charges which may arise in connection with a warranty claim.

If the pump fails during the warranty period, after confirmation by our technical department, we will provide spare parts free of charge. Customers will need to bear the shipping cost.

Exceptions:

- The warranty shall not apply to repairs or service necessitated by normal wear and tear or for lack of reasonable and proper maintenance.

- All tubing and pumping accessories as consumable items are excluded.

- Electrical surge as a cause of failure is excluded.

- Chemical attack is excluded.

- Improper operation or man-made damage as a cause of failure is excluded.